## Slide 1: CV app
3/14/2019
Guillaume von Gastrow

## Slide 2: CV app measurement cycle
time
Temperature
CV
measurement
meas_flag set to 1
meas_flag set to 0
meas_flag set to 1
Start bias if:
stressing temperature is reached
meas_flag=1
If
steady-state bias time is over
hotplate still at stressing temperature
Then: turn off bias, set T to Troom, turn on fan, set fan_flag to 1
Record starting time of the steady-state bias
Start measurement if:
Steady-state bias time is over
Hotplate at room temperature
meas_flag=0
CV
measurement
bias
After each pin measurement, check on other hotplates if:
Bias ready to be turned on
Bias ready to be turned off
Fan ready to be turned off
Turn off fan if:
Fan_flag=1
Cooling temperature reached
Then set fan_flag to 0
stress_completed_flag set to 0
stress_completed_flag set to 1
Notes: meas_flag is used during the steady-state bias to avoid turning on bias if it is already on.
Problem with the current implementation of the way the fan is turned off: will be constantly calling the Arduino to turn it off as long as the hotplate is at room temperature.

## Slide 3: CV app measurement cycle (actual parameter names)
time
Temperature
CV
measurement
meas_flag set to 1
meas_flag set to 0
meas_flag set to 1
Start bias if:
T=stressT
meas_flag=1
Stop bias if
t-tstartproc  setbiastime
Record tstartproc
Set T = coolT
Turn on the fan and set fanflag to 1
Start measurement if:
t-tstartproc  setbiasbiastime
T=coolT
meas_flag=0
CV
measurement
bias

## Slide 4: CV app data logs (T, time, I)
time
Temperature
CV
measurement
meas_flag set to 1
meas_flag set to 1
CV
measurement
bias
Log after each CV measurement
Log during temperature rampup
Log during stress
If time elapsed since last log is larger than the log step
Log during temperature rampdown
Log when at room temperature before the next measurement
In the current 4-pin app, the time interval between temperature logging and plotting is :
10 s during the initial temperature cooldown
User-defined during temperature rampup: WtS (5 s by default, usually set to 30 s)
User-defined during temperature cooldown: WtC (5 s by default, usually set to 30 s)
User-defined and plotted at the same time as current during bias (defined by the number of current points)
Record starting time of the steady-state bias
Notes: meas_flag is used during the steady-state bias to avoid turning on bias if it is already on.
Problem with the current implementation of the way the fan is turned off: will be constantly calling the Arduino to turn it off as long as the hotplate is at room temperature.

## Slide 5: CV app arguments and data structure – Local variables
Main arguments
All parameters are stored in non-scalar structure arrays to allow access for each measurement unit (MU).
MD structure (measurement data) contains measurement data and parameters
HW structure (hardware) contains the Arduino, thermocouple and hotplate handles, and visa objects to Keithley and impedance analyzer
HW structure
Arduino contains the Arduino handles for each hotplate:
HW(1).Arduino, HW(2).Arduino, HW(3).Arduino
Therm contains the Arduino connection to the thermocouple for each hotplate
HW(1).Therm, HW(2).Therm, HW(3).Therm
HP contains the hotplate handles
HW(1).HP, HW(2).HP, HW(3).HP
IMPA contains the visa object of the impedance analyzer (stored in MU1 because same for all MUs)
HW(1).IMPA
KEITH contains the visa object to the Keithley (stored in MU1 because same for all MUs)
HW(1).KEITH

## Slide 6: CV app arguments and data structure – Local variables
Main arguments (continued)
MD structure (Measurement Data)
PinState structure contains the on/off sate of all pins (off in case one pin is not used):
MD(mu).PinState=[1 1 1 1 1 1 1 1], MD(2).PinState, MD(3).PinState.
ArdP contains the Arduino pin numbers connected to the relays:
MD(mu).ArdP=[2 3 4 5 6 7 8], MD(2).ArdP, MD(3).ArdP
MDdata contains flags, times
MD(mu).MDdata.startbiastime=[300, 4000, 8500] bias starting time (recorded by the system)
MD(mu).MDdata.meas_flag=0 measurement flag value (recorded by the system)
MD(MUnb).MDdata.stress_completed_flag flag to indicate whether a stress cycle has been completed
MD(mu).MDdata.CVStartTime=[2000, 5000, 8000] CV measurement starting time at each iteration (recorded)
MD(mu).MDdata.fanflag flag indicating whether fan is on or off
MD(1).MDdata.Err Allowed error in temperature to start a measurement (hard-coded in StartProc)
MD(mu).MDdata.finish_flag Flag indicating that measurements are finished for one hotplate
MD(mu).MDdata.cycle_counter Number of cycles performed (= number of bias iterations)

## Slide 7
Main arguments (continued)
MD structure (Measurement Data)
Plots contains the figure handles for each pin (need to be defined at beginning of StartProc)
MD(mu).Plots.CV Object array that contains the handles to the CV plots for each pin of Measurement Unit n
MD(mu).Plots.CVby2 Object array that contains the handles to the CV derivative plots for each pin of MU n
MD(mu).Plots.Temp Object array that contains the temperature plot handle
MD(mu).Plots.Current Object array that contains the plot handle of the current measured by the Keithley for all pins
MD(mu).Plots.VfbTime Object array that contains the flatband vs time plot handle

## Slide 8
Main arguments (continued)
MD structure (Measurement Data)
ExpData contains the experimentally measured data and log data
MD(mu).ExpData.Pin(i).C matrix containing the capacitance measurement data for hotplate n and pin i
MD(mu).ExpData.Pin(i).V matrix containing the voltage measurement data for hotplate n and pin I
MD(mu).ExpData.Pin(i).R matrix containing the resistance measurement data for hotplate n and pin I
MD(mu).ExpData.Pin(i).Vinput vector containing the user-defined voltage parameters on the app panel
MD(mu).ExpData.Pin(i).tfb vector containing the calculated times (not measured) between each CV
MD(mu).ExpData.Pin(i).Vfb vector containing the extracted flatbands measured for each iteration
MD(mu).ExpData.Pin(i).VfbAve vector containing the flatband at each time point, averaged over the number of iterations
MD(mu).ExpData.Pin(i).Cfit vector containing selected capacitance data (noisy part removed)
MD(mu).ExpData.Pin(i).Vfit vector containing selected voltage data (noisy part removed)
MD(mu).ExpData.Pin(i).VfbStd vector containing the standard deviation of flatband at each time point

## Slide 9
Main arguments (continued)
MD structure (Measurement Data)
ExpData contains the experimentally measured data and log data
MD(mu).ExpData.Setup.PreBiasTime
MD(mu).ExpData.Setup.t_offset_unit unit of the time offset (to define in startproc)
MD(mu).ExpData.Setup.t_offset_value unit of the time offset (to define in startproc)
MD(mu).ExpData.Setup.stressBiasValue value of steady-state bias (needs to be the same for all hotplates)
MD(mu).ExpData.Setup.biastime_sec steady-state bias time converted to s (defined in startproc)
MD(mu).ExpData.Setup.TempC cooling temperature (define in startproc)
MD(mu).ExpData.Setup.TempH stress temperature (define in startproc)
MD(mu).ExpData.Setup.CalTempH stress temperature corrected with calibration curve (defined in startproc)
MD(mu).ExpData.Setup.CalTempC cool temperature corrected with calibration curve (defined in startproc)

## Slide 10
Main arguments (continued)
MD structure (Measurement Data)
ExpData contains the experimentally measured data and log data
MD(mu).ExpData.Setup.IterM Number of repeated measurements per iteration
MD(mu).ExpData.log.T Array of temperature log (different for each MU) TO INITIALIZE
MD(mu).ExpData.log.Ttime Array of time log (time at which temperature is acquired) TO INITIALIZE
MD(mu).ExpData.log.Vbias=[0 0 0 7 7 7 7 7 7 7] value of bias voltage log (saved at the same time as t)
MD(1).ExpData.log.I value of the current log from Keithley
MD(1).ExpData.log.Itime Time at which current is acquired. Current is acquired independently from hotplate temperature, so need to be stored separately
MD(mu).ExpData.log.Pin(i).tCV vector containing the time at which each CV was measured (NEED TO ADD!)
MD(mu).ExpData.log. startbiastime vector containing the time at which bias was started on each unit (NEED TO ADD!). It contains the same values as MD(mu).MDdata.startbiastime.